GLOBALEMED, LLC.

 1101 King Street   Suite 370  Alexandria, VA  22314

Telephone Number:  (703) 894-0710    Fax Number: (703) 894-0725

STABILITY AND QUALITY CONTROL TESTS FOR THE HIV 1&2 RAPID DEVICE (SMART CHECK HIV 1&2)












                           PAGE 1/1

HIV 1&2 RAPID TESTS (Smart Check HIV 1&2)

Fifty rapid test were stored at three different temperatures for three months to test their specificity and sensitivity:

Stored at 4’C for three months.  Lot#: 09071585, Exp. Date: 09/01
	HIV 1&2 Rapid Test
	Patient Positive to HIV 1&2
	Patient Positive to HCV
	Patient Positive to HBV

	POSITIVE  RESULTS
	50
	0
	0

	NEGATIVE RESULTS
	0
	0
	0


Sample number=  50;   Specificity=  100%, Sensitivity= 100%

Stored at Room temperature (12’C-25’C) for three months (usual way of storage) . Lot#: 09071585, Exp. Date: 09/01
	HIV 1&2 Rapid Test
	Patient Positive to HIV 1&2
	Patient Positive to HCV
	Patient Positive to HBV

	POSITIVE RESULTS
	50
	0
	0

	NEGATIVE RESULTS
	0
	0
	0



Sample number=  50;   Specificity=  100%, Sensitivity= 100%

Stored at 37’C for three months. Lot#: 09071585, Exp. Date: 09/01
	HIV 1&2 Rapid Test
	Patient Positive to HIV 1&2
	Patient Positive to HCV
	Patient Positive to HBV

	POSITIVE RESULTS
	40
	0
	0

	NEGATIVE RESULTS
	5
	0
	0


Sample number=  50  Specificity=  100%, Sensitivity= 80%








                                                                                                GLOBALEMED, LLC.

                  






1101 King Street, Suite 370

Alexandria, VA  22314 USA

QUALITY CONTROL





June 15th, 2000. Analysis done by: Elina Riesgo, M.T.














